## Slide 1
Bangladesh University of Engineering and Technology
BME 300 : Biomedical Engineering Design 1
Mir Maheen Labib Mehzabin Morshed Raka Stu ID: 201518021 Stu ID:201518022

## Slide 2: Prototype of a Portable Hemodialysis Machine

## Slide 3: Hemodialysis
Hemodialysis is a method for removing waste products such as urea, potassium as well as free water from blood when the kidneys are incapable of this . People who are suffering from renal failure need hemodialysis.

## Slide 4: Hemodialysis Machine

## Slide 5: Problems leading to our prototype:
The size of the machine
Immobility of the machine
Fixed water treatment plant
Can be used only in the dialysis unit of the hospital
Power supply
Too costly

## Slide 6: Border impact of our device:
mobility
Usefulness in any unit of the hospital
Battery Power supply
Cost effectiveness

## Slide 7: Factors to final design:
                Conform to target excretion rates
                Maintains blood pressure
                Waste is excreted from the system in a sanitary manner
                Stable saline stand
                Proper alarm system
                Feedback flow control

## Slide 8: Fabrication process: Materials
Structure and Casing:
1. PVC (Polyvinylchloride) plastic rack
High UTS (52 MPa ), Low relative density
(1.42-1.48 ), Sufficient Hardness (80)
Good impact strength (20 kj/(20◦ C))
2. Polymeric PVC sheet (5 and 12 mm height)
for casing and additional stability

## Slide 9: Fabrication process: Materials
Fluid Flow line:
0.8 mm (1/4 inch) diameter clear PVC plastic pipe for DI water and Dialysate solution flow,clear visibility and flexibility (50-80% Elongation at brake)
Rubber connector high elasticity (up to 700% elongation)
Blood Flow Line:
Highly biocompatible one time use rubber pipe hemodialysis blood line sets manufactured specially for dialysis purpose

## Slide 10: Fabrication process: Materials
Peristaltic Pump:
12V NEMA 23 stepper motor
80 mm diameter Mild Steel Bar pump head, casing and rotter high yield strength (370 MPa) modulus of elasticity (205 GPa)
8/22/7 Deep Groove Ball Bearing low friction

## Slide 11: Fabrication process: Materials:
Syringe Pump:
1. MG996R servo motor:
5 ml syringe
Metal wires normal pulley system
Water Pumps: 12V DC water pumps
Mixing chamber:
PVC plastic container

## Slide 12: Fabrication process: Materials:

## Slide 13: Fabrication process: Materials:

## Slide 14: Fabrication process: Materials:

## Slide 15: Fabrication process: Methods: dialysate flow line & controller

## Slide 16: Fabrication process: Methods: blood flow line & controller

## Slide 17: Fabrication process: Methods:
Power connection:

## Slide 18: PRODUCT DESIGN SPECIFICATION:
Blood Flow rate= 200 to 220 ml/min
Heparin used= 1 cc
Saline = 1 bag or 450 ml
Ultrafiltration goal= 1000ml/ 4 hours (dialysate A and B)
Ultrafiltration rate= 250 ml/hour (dialysate A and B)
Dialysate solution flow rate = 550 ml/min (usually)
DI water flow rate =500 ml/min (usually)
Temperature monitoring= 37◦C constant
Device should be reduced in size and weight.
The power supply should be mobile for Portable dialysis machine.

## Slide 19: Evaluation: dialysate solution flow rate
Chart: Flow rate vs Duty cycle
| 100 | 80 | 60 | 45 | 40 | 30
Flow rate | 2200 | 1714.28 | 1440 | 1058 | 900 | 692

## Slide 20: Evaluation: dialysate A & B flow rate
Chart: Flow rate vs Tube Openning
| 100 | 80 | 60 | 50 | 40 | 20
Flow rate | 170 | 122.95 | 67.8 | 46.5 | 33 | 16.84

## Slide 21: Evaluation:
Dialysate Pressure result:
Hagen-Poiseuille equation
Where:
ΔP is the pressure difference between the two ends,
L is the length of pipe,
μ is the dynamic viscosity,
Q is the volumetric flow rate,
R is the pipe radius.

## Slide 22: Evaluation:
Dialysate Pressure result: CONT
tube average diameter = 0.006 m.
Average length of the tube =1.52 meter.
The dynamic viscosity of DI water = 8.90× Pa.s or 0.89 cp at 25◦C temperature.
Flow rate = 690 ml/min.
The approximate pressure difference value is 490 Pa with respect to the atmospheric pressure.

## Slide 23: Evaluation:
Peristaltic Pump Testing:
1. 2 shoe for compressing the tube
Height of the bearing =7 mm
which was lower than tube diameter= 10 mm.
The minimal distance between the bearing and casing wall = 4 mm instead of 2.5 mm due to the machine cutting error of metals.
All these reasons made high amount of backflow of fluid and reducing the flow rate to 1 to 2 ml/min only.

## Slide 24: Evaluation:
Air Detector Testing:
Photodiode: An LED and the diode were placed across the blood tube in a dark medium,the differential value between air or fluid presence= 2 to 4 analog value range.
Sonar sensor: 2 to 3 microsecond pulse Insufficient for blood tube penetration.

## Slide 25: Evaluation:
Air Detector Testing: PIR sensor:
It gave digital output value at a continuous rate when the blood in present. If air was present in the tube it returned no value for long time.

## Slide 26: Evaluation:
Heater Testing:
6 watt at 12V supply.
Temperature change =0.72◦C/min.
Temperature maintaining 37◦C.
Power Line Testing:
The voltage output =12.3V (connected to the battery at full charge)
The voltage output =13.8v (connected to the 220V AC line)
A voltage limiting POT used all the 5V supplies were made using 7805 IC

## Slide 27: Evaluation:
Heparin Pump Testing:
syringe total capacity = 5 ml
The maximum volume it could deliver =2.5 ml
due to mechanical system loss at a short 30 min period of time.
Sensor Testing:
Flow Sensor Calibration: the calibration factor was set to 4.5
Pressure Sensor Testing: 0.5 m water level testing the calibration factor was set to 4.5

## Slide 28: Specific info. Of final prototype:
Peristaltic Pump:
Designed Pump Clinically used pump
Casing
Rotor
Bearing
Motor

## Slide 29: Specific info. Of final prototype
Dialysate flow line:
Pump
Pipes
Flow meter
Controller
circuit

## Slide 30: Specific info. Of final prototype
Power line and mixing chamber:
Mixing chamber
Battery
Relay board
Adapter
Switches

## Slide 31: Specific info. Of final prototype
AV line controller circuits and sensors:
AV line controller
LCD
L298n motor driver
POT
Pressure sensor

## Slide 32: Specific info. Of final prototype
Interior: Exterior:

## Slide 33: Budget
Total Cost:
Continued…

## Slide 34: Budget

## Slide 35: Have we found the solution?
Size reduced
Mobility increased
Can be used in other units of a hospital as well as at home
Cost will eventually reduce
Power & water treatment problem will hopefully be addressed later on

## Slide 36: Limitations:
Manual dialysate flow control
Faulty bubble detector
Small range syringe pump
Absence of cooler
Peristaltic pump
Absence of folding saline stand

## Slide 37: Future goals:
Control dialysate flow by pumps
Construct perfect bubble detecting system
Making full range syringe pump
Developing cooling system
Peristaltic pump set up
Making folding saline stand

## Slide 39
Thank you!